## OREAS standard values
620 | Ag, Silver (ppm) | Al, Aluminium (wt.%) | As, Arsenic (ppm) | Ba, Barium (ppm) | Bi, Bismuth (ppm) | Ca, Calcium (wt.%) | Cd, Cadmium (ppm) | Ce, Cerium (ppm) | Co, Cobalt (ppm) | Cr, Chromium (ppm) | Cs, Cesium (ppm) | Cu, Copper (wt.%) | Fe, Iron (wt.%) | Ga, Gallium (ppm) | In, Indium (ppm) | K, Potassium (wt.%) | La, Lanthanum (ppm) | Li, Lithium (ppm) | Mg, Magnesium (wt.%) | Mn, Manganese (ppm) | Mo, Molybdenum (ppm) | Nb, Niobium (ppm) | Nd, Neodymium (ppm) | P, Phosphorus (ppm) | Pb, Lead (wt.%) | Pr, Praseodymium (ppm) | Rb, Rubidium (ppm) | S, Sulphur (wt.%) | Sb, Antimony (ppm) | Si, Silicon (wt.%) | Sr, Strontium (ppm) | Ta, Tantalum (ppm) | Th, Thorium (ppm) | Ti, Titanium (wt.%) | Tl, Thallium (ppm) | U, Uranium (ppm) | V, Vanadium (ppm) | W, Tungsten (ppm) | Y, Yttrium (ppm) | Yb, Ytterbium (ppm) | Zn, Zinc (wt.%)
| 40.0397 | 7.0582 | 53.703 | 2754.1817 | 1.9392 | 1.6319 | 166.9269 | 68.6927 | 13.3707 | 25.4856 | 5.4483 | 0.17566 | 3.0128 | 24.1764 | 1.2358 | 2.6986 | 35.8917 | 20.679 | 0.34751 | 449.2646 | 10.5379 | 15.165 | 33.3037 | 379.6667 | 0.77175 | 8.9644 | 122.534 | 2.49 | 81.0453 | 29.8177 | 141.8969 | 1.1642 | 11.1976 | 0.15499 | 1.5615 | 4.0575 | 26.8747 | 2.2581 | 15.4585 | 0.88367 | 3.1408
621 | Ag, Silver (ppm) | Al, Aluminium (wt.%) | As, Arsenic (ppm) | Ba, Barium (ppm) | Bi, Bismuth (ppm) | Ca, Calcium (wt.%) | Cd, Cadmium (ppm) | Ce, Cerium (ppm) | Co, Cobalt (ppm) | Cr, Chromium (ppm) | Cs, Cesium (ppm) | Cu, Copper (wt.%) | Fe, Iron (wt.%) | Ga, Gallium (ppm) | In, Indium (ppm) | K, Potassium (wt.%) | La, Lanthanum (ppm) | Mg, Magnesium (wt.%) | Mn, Manganese (ppm) | Mo, Molybdenum (ppm) | Nb, Niobium (ppm) | Nd, Neodymium (ppm) | P, Phosphorus (ppm) | Pb, Lead (wt.%) | Pr, Praseodymium (ppm) | Rb, Rubidium (ppm) | S, Sulphur (wt.%) | Sb, Antimony (ppm) | Si, Silicon (wt.%) | Sr, Strontium (ppm) | Th, Thorium (ppm) | Ti, Titanium (wt.%) | Tl, Thallium (ppm) | U, Uranium (ppm) | V, Vanadium (ppm) | W, Tungsten (ppm) | Y, Yttrium (ppm) | Yb, Ytterbium (ppm) | Zn, Zinc (wt.%) |  |  |  |  | Be, Beryllium (ppm)
| 68.5441 | 6.632 | 84.674 | 2612.4133 | 3.9965 | 2.0041 | 295.2899 | 51.6393 | 31.42 | 48.7099 | 3.5945 | 0.36821 | 3.7134 | 26.5359 | 1.9327 | 2.2319 | 26.0967 | 0.51642 | 554.0139 | 13.5105 | 10.4 | 24.2363 | 392.8333 | 1.3254 | 6.636 | 88.7017 | 4.506 | 145.8245 | 28.0508 | 101.2814 | 8.5626 | 0.18101 | 1.9867 | 2.9999 | 36.2778 | 2.6327 | 13.9299 | 1.034 | 5.2174 |  |  |  |  | 1.9973
622 | Ag, Silver (ppm) | Al, Aluminium (wt.%) | As, Arsenic (ppm) | Ba, Barium (ppm) | Bi, Bismuth (ppm) | Ca, Calcium (wt.%) | Cd, Cadmium (ppm) | Ce, Cerium (ppm) | Co, Cobalt (ppm) | Cr, Chromium (ppm) | Cs, Cesium (ppm) | Cu, Copper (wt.%) | Fe, Iron (wt.%) | Ga, Gallium (ppm) | In, Indium (ppm) | K, Potassium (wt.%) | La, Lanthanum (ppm) | Li, Lithium (ppm) | Mg, Magnesium (wt.%) | Mn, Manganese (wt.%) | Mo, Molybdenum (ppm) | Nb, Niobium (ppm) | Nd, Neodymium (ppm) | Ni, Nickel (ppm) | Pb, Lead (wt.%) | Pr, Praseodymium (ppm) | Rb, Rubidium (ppm) | S, Sulphur (wt.%) | Sb, Antimony (ppm) | Sc, Scandium (ppm) | Si, Silicon (wt.%) | Sn, Tin (ppm) | Sr, Strontium (ppm) | Th, Thorium (ppm) | Ti, Titanium (wt.%) | Tl, Thallium (ppm) | U, Uranium (ppm) | V, Vanadium (ppm) | W, Tungsten (ppm) | Y, Yttrium (ppm) | Yb, Ytterbium (ppm) | Zn, Zinc (wt.%)
| 102.6817 | 5.8258 | 110.3137 | 2311.3356 | 4.9601 | 2.1368 | 454.2568 | 34.8013 | 37.6831 | 48.0615 | 1.8826 | 0.48367 | 4.2585 | 26.8547 | 4.8531 | 1.7826 | 18.798 | 9.5634 | 0.57068 | 0.061277 | 19.4495 | 6.2359 | 15.7233 | < 40 | 2.1053 | 4.2936 | 57.5778 | 7.6778 | 197.7355 | < 10 | 25.0869 | 61.6742 | 63.2444 | 5.4722 | 0.18723 | 3.6238 | 1.5527 | 45.3714 | 2.9288 | 12.3657 | 1.1687 | 9.8985
623 | Ag, Silver (ppm) | Al, Aluminium (wt.%) | As, Arsenic (ppm) | Ba, Barium (ppm) | Bi, Bismuth (ppm) | Ca, Calcium (wt.%) | Cd, Cadmium (ppm) | Ce, Cerium (ppm) | Co, Cobalt (ppm) | Cr, Chromium (ppm) | Cs, Cesium (ppm) | Cu, Copper (wt.%) | Fe, Iron (wt.%) | Ga, Gallium (ppm) | In, Indium (ppm) | K, Potassium (wt.%) | La, Lanthanum (ppm) | Li, Lithium (ppm) | Mg, Magnesium (wt.%) | Mn, Manganese (wt.%) | Mo, Molybdenum (ppm) | Nb, Niobium (ppm) | Nd, Neodymium (ppm) | Ni, Nickel (ppm) | Pb, Lead (wt.%) | Pr, Praseodymium (ppm) | Rb, Rubidium (ppm) | S, Sulphur (wt.%) | Sb, Antimony (ppm) | Sc, Scandium (ppm) | Si, Silicon (wt.%) | Sr, Strontium (ppm) | Th, Thorium (ppm) | Ti, Titanium (wt.%) | Tl, Thallium (ppm) | U, Uranium (ppm) | W, Tungsten (ppm) | Y, Yttrium (ppm) | Yb, Ytterbium (ppm) | Zn, Zinc (wt.%)
| 22.6512 | 5.2498 | 86.0053 | 1394.5826 | 17.6968 | 1.3699 | 54.598 | 49.8907 | 225.1316 | 29.9566 | 3.0016 | 1.7265 | 13.447 | 23.6985 | 2.0445 | 1.5264 | 24.8622 | 17.9736 | 1.2294 | 0.060477 | 10.2045 | 10.4374 | 24.8883 | < 50 | 0.24727 | 6.5298 | 66.6862 | 8.8333 | 27.9668 | < 10 | 24.4529 | 89.9114 | 7.2896 | 0.15009 | 0.70858 | 2.779 | 4.223 | 19.14 | 1.7553 | 1.0348
624 | Ag, Silver (ppm) | Al, Aluminium (wt.%) | As, Arsenic (ppm) | Ba, Barium (ppm) | Bi, Bismuth (ppm) | Ca, Calcium (wt.%) | Cd, Cadmium (ppm) | Ce, Cerium (ppm) | Co, Cobalt (ppm) | Cr, Chromium (ppm) | Cs, Cesium (ppm) | Cu, Copper (wt.%) | Fe, Iron (wt.%) | Ga, Gallium (ppm) | In, Indium (ppm) | K, Potassium (wt.%) | La, Lanthanum (ppm) | Li, Lithium (ppm) | Mg, Magnesium (wt.%) | Mn, Manganese (wt.%) | Mo, Molybdenum (ppm) | Nb, Niobium (ppm) | Nd, Neodymium (ppm) | Ni, Nickel (ppm) | P, Phosphorus (wt.%) | Pb, Lead (wt.%) | Pr, Praseodymium (ppm) | Rb, Rubidium (ppm) | S, Sulphur (wt.%) | Sb, Antimony (ppm) | Sc, Scandium (ppm) | Si, Silicon (wt.%) | Sn, Tin (ppm) | Sr, Strontium (ppm) | Th, Thorium (ppm) | Ti, Titanium (wt.%) | Tl, Thallium (ppm) | U, Uranium (ppm) | V, Vanadium (ppm) | W, Tungsten (ppm) | Y, Yttrium (ppm) | Yb, Ytterbium (ppm) | Zn, Zinc (wt.%)
| 46.1153 | 4.3237 | 114.5497 | 1070.044 | 21.2948 | 1.4886 | 132.5727 | 32.9293 | 273.1139 | < 50 | 1.3203 | 3.0794 | 16.3109 | 22.1071 | 4.1434 | 0.99101 | 17.3197 | 10.2716 | 1.3111 | 0.065984 | 17.8308 | 5.775 | 16.8277 | < 20 | < 0.1 | 0.61199 | 4.2742 | 33.0439 | 13.1789 | 71.593 | < 10 | 20.4739 | < 50 | 47.6023 | 4.1245 | 0.14567 | 0.94401 | 1.3398 | 43.2709 | 4.584 | 17.2667 | 1.9393 | 2.4129

## pXRF median measured values
OREAS no. | Nd | Nd | Pr | Pr | Ce | Ce | La | La | Ba | Ba | Y | Y | P | P | S | S | Cl | Cl | K | K | Ca | Ca | Ti | Ti | V | V | Cr | Cr | Mn | Mn | Fe | Fe | Co | Co | Ni | Ni | Cu | Cu | Zn | Zn | As | As | Se | Se | Rb | Rb | Sr | Sr | Zr | Zr | Nb | Nb | Mo | Mo | Ag | Ag | Cd | Cd | Sn | Sn | Sb | Sb | Ta | Ta | W | W | Au | Au | Hg | Hg | Pb | Pb | Bi | Bi | Th | Th | U | U
OREAS 620 | <LOD | 310 | <LOD | 168 | <LOD | 91 | <LOD | 71 | 2812 | 30 | 93 | 5 | 9471 | 2418 | 13133 | 327 | 1247 | 131 | 25263 | 332 | 16801 | 217 | 2562 | 43 | 204 | 5 | 85 | 5 | 508 | 10 | 39632 | 258 | <LOD | 117 | <LOD | 20 | 1844 | 27 | 36628 | 220 | 97 | 17 | <LOD | 6 | 122 | 2 | 146 | 3 | 202 | 4 | 8.3 | 1.3 | 8.8 | 1.4 | 39.4 | 1.7 | 157 | 3 | 41 | 3 | 90 | 4 | 115 | 10 | <LOD | 155 | <LOD | 20 | <LOD | 24 | 7862 | 51 | <LOD | 29 | 19 | 2 | 20 | 3
OREAS 620 | <LOD | 312 | <LOD | 170 | <LOD | 93 | <LOD | 72 | 2835 | 30 | 89 | 5 | 10901 | 2456 | 12416 | 320 | 1338 | 133 | 25351 | 335 | 16953 | 220 | 2546 | 44 | 201 | 5 | 87 | 5 | 503 | 10 | 39313 | 254 | <LOD | 116 | <LOD | 21 | 1894 | 26 | 36516 | 217 | 60 | 16 | <LOD | 6 | 119.2 | 2 | 154 | 3 | 211 | 4 | 11.4 | 1.3 | 5.9 | 1.4 | 38.6 | 1.7 | 158 | 3 | 36 | 3 | 91 | 4 | 89 | 10 | <LOD | 152 | <LOD | 20 | <LOD | 23 | 7834 | 50 | <LOD | 29 | 17 | 2 | 23 | 3
OREAS 620 | <LOD | 314 | <LOD | 171 | <LOD | 93 | <LOD | 73 | 2735 | 30 | 93 | 5 | 8002 | 2371 | 12072 | 316 | 1036 | 127 | 24336 | 326 | 16138 | 212 | 2454 | 43 | 196 | 5 | 84 | 5 | 469 | 10 | 38683 | 258 | <LOD | 118 | <LOD | 21 | 1894 | 27 | 35874 | 219 | 74 | 17 | <LOD | 6 | 116 | 2 | 146 | 3 | 204 | 4 | 9.8 | 1.4 | 9.1 | 1.4 | 40.8 | 1.7 | 152 | 3 | 36 | 3 | 89 | 4 | 63 | 10 | 160 | 52 | <LOD | 20 | <LOD | 24 | 7663 | 51 | <LOD | 30 | 17 | 2 | 21 | 3
OREAS 620 | <LOD | 317 | <LOD | 171 | <LOD | 93 | <LOD | 73 | 2757 | 30 | 99 | 5 | 8549 | 2394 | 12016 | 316 | 1397 | 134 | 25094 | 334 | 16249 | 214 | 2463 | 43 | 207 | 5 | 77 | 5 | 504 | 11 | 38921 | 258 | <LOD | 118 | <LOD | 22 | 2009 | 27 | 36209 | 220 | 131 | 17 | <LOD | 6 | 118 | 2 | 144 | 3 | 204 | 4 | 11.3 | 1.4 | 5.6 | 1.4 | 38.5 | 1.7 | 153 | 3 | 39 | 3 | 92 | 4 | <LOD | 29 | <LOD | 155 | <LOD | 20 | <LOD | 24 | 7670 | 51 | <LOD | 30 | 18 | 2 | 22 | 3
OREAS 620 | <LOD | 314 | <LOD | 169 | 121 | 31 | <LOD | 72 | 2823 | 30 | 97 | 5 | 9150 | 2375 | 12058 | 312 | 1057 | 125 | 24539 | 324 | 16354 | 212 | 2492 | 43 | 197 | 5 | 80 | 5 | 487 | 10 | 39454 | 257 | <LOD | 116 | <LOD | 21 | 2023 | 26 | 36550 | 219 | 70 | 17 | <LOD | 6 | 121 | 2 | 148 | 3 | 212 | 4 | 10.5 | 1.3 | 6.7 | 1.4 | 38.9 | 1.7 | 154 | 3 | 43 | 3 | 95 | 4 | <LOD | 29 | 289 | 51 | <LOD | 20 | <LOD | 24 | 7805 | 50 | <LOD | 29 | 14.9 | 2 | 16 | 3
OREAS 620 | <LOD | 309 | <LOD | 168 | <LOD | 92 | <LOD | 71 | 2757 | 29 | 85 | 5 | 13732 | 2512 | 12689 | 323 | 1217 | 131 | 25065 | 332 | 16951 | 219 | 2583 | 44 | 202 | 5 | 77 | 5 | 492 | 10 | 39866 | 261 | <LOD | 118 | <LOD | 21 | 1903 | 27 | 36944 | 223 | 82 | 17 | <LOD | 6 | 121 | 2 | 149 | 3 | 209 | 4 | 12.3 | 1.4 | 8.9 | 1.4 | 37.2 | 1.7 | 152 | 3 | 32 | 3 | 92 | 4 | 91 | 10 | <LOD | 156 | <LOD | 20 | <LOD | 24 | 7879 | 51 | <LOD | 30 | 19 | 2 | 16 | 3
OREAS 620 | <LOD | 311 | <LOD | 169 | <LOD | 92 | <LOD | 72 | 2746 | 30 | 92 | 5 | 10053 | 2453 | 12048 | 319 | 1031 | 128 | 24619 | 333 | 16485 | 218 | 2482 | 43 | 195 | 5 | 85 | 5 | 479 | 10 | 39164 | 256 | <LOD | 117 | <LOD | 21 | 1905 | 26 | 36346 | 218 | 134 | 17 | <LOD | 6 | 121 | 2 | 155 | 3 | 208 | 4 | 10.7 | 1.3 | 5.9 | 1.4 | 40.1 | 1.7 | 153 | 3 | 34 | 3 | 87 | 4 | 61 | 10 | <LOD | 154 | <LOD | 20 | <LOD | 24 | 7708 | 50 | <LOD | 29 | 22 | 2 | 23 | 3
OREAS 620 | <LOD | 313 | <LOD | 169 | <LOD | 92 | <LOD | 72 | 2817 | 30 | 101 | 5 | 9300 | 2417 | 12936 | 326 | 965 | 125 | 25162 | 332 | 16784 | 217 | 2556 | 43 | 200 | 5 | 79 | 5 | 490 | 10 | 39827 | 261 | <LOD | 119 | <LOD | 21 | 1988 | 27 | 36742 | 222 | 108 | 17 | <LOD | 6 | 119 | 2 | 150 | 3 | 217 | 4 | 11.5 | 1.4 | 4.6 | 1.4 | 37.7 | 1.7 | 153 | 3 | 39 | 3 | 95 | 4 | <LOD | 29 | 217 | 52 | <LOD | 20 | <LOD | 24 | 7837 | 51 | <LOD | 30 | 13.4 | 2 | 29 | 4
Oreas 620 | <LOD | 309 | <LOD | 168 | <LOD | 92 | <LOD | 71 | 2801 | 30 | 86 | 5 | 10858 | 2774 | 14391 | 377 | 1287 | 149 | 28964 | 403 | 19076 | 261 | 2915 | 52 | 234 | 6 | 108 | 6 | 565 | 12 | 39012 | 255 | <LOD | 117 | <LOD | 21 | 1865 | 26 | 36366 | 218 | 68 | 17 | <LOD | 6 | 116.7 | 2 | 148 | 3 | 208 | 4 | 11 | 1.3 | 6.6 | 1.4 | 33.8 | 1.7 | 150 | 3 | 37 | 3 | 95 | 4 | 100 | 10 | <LOD | 154 | <LOD | 20 | <LOD | 24 | 7804 | 51 | <LOD | 29 | 15 | 2 | 18 | 3
OREAS 620 | <LOD | 314 | <LOD | 170 | <LOD | 94 | <LOD | 73 | 2834 | 30 | 92 | 5 | 12462 | 2828 | 14414 | 378 | 980 | 143 | 28852 | 403 | 19548 | 267 | 2942 | 52 | 229 | 6 | 97 | 6 | 563 | 12 | 40088 | 266 | <LOD | 120 | <LOD | 22 | 1921 | 27 | 37384 | 228 | 111 | 17 | <LOD | 6 | 117 | 2 | 146 | 3 | 207 | 4 | 12.9 | 1.4 | 9.2 | 1.4 | 37.6 | 1.7 | 155 | 3 | 41 | 3 | 87 | 4 | 89 | 10 | <LOD | 159 | <LOD | 21 | <LOD | 24 | 7949 | 52 | <LOD | 30 | 16 | 2 | 20 | 3
Oreas 620 | <LOD | 310 | <LOD | 169 | <LOD | 92 | <LOD | 72 | 2787 | 30 | 97 | 5 | 7328 | 2388 | 12703 | 325 | 1231 | 132 | 25536 | 338 | 16743 | 218 | 2649 | 45 | 198 | 5 | 90 | 5 | 506 | 11 | 39377 | 259 | <LOD | 118 | <LOD | 22 | 1856 | 27 | 36567 | 221 | 115 | 17 | <LOD | 6 | 117 | 2 | 149 | 3 | 214 | 4 | 12.5 | 1.4 | 7.5 | 1.4 | 36.6 | 1.7 | 155 | 3 | 39 | 3 | 89 | 4 | 94 | 10 | <LOD | 155 | <LOD | 20 | <LOD | 24 | 7784 | 51 | <LOD | 30 | 17 | 2 | 19 | 3
OREAS 620 | <LOD | 320 | <LOD | 172 | 114 | 32 | <LOD | 73 | 2814 | 30 | 95 | 5 | 10570 | 2468 | 12730 | 326 | 1170 | 131 | 25117 | 335 | 16977 | 221 | 2578 | 44 | 203 | 5 | 85 | 5 | 498 | 10 | 39532 | 260 | <LOD | 118 | <LOD | 21 | 1935 | 27 | 36595 | 222 | 100 | 17 | <LOD | 6 | 117 | 2 | 149 | 3 | 203 | 4 | 12.1 | 1.4 | 8.3 | 1.4 | 40 | 1.7 | 155 | 3 | 44 | 3 | 94 | 4 | 54 | 10 | <LOD | 156 | <LOD | 21 | <LOD | 24 | 7800 | 51 | <LOD | 30 | 17 | 2 | 19 | 3
OREAS 620 <3 | <LOD | 311 | <LOD | 168 | <LOD | 92 | <LOD | 71 | 2740 | 29 | 97 | 5 | 7642 | 2337 | 12377 | 315 | 1106 | 126 | 24815 | 325 | 16534 | 212 | 2453 | 42 | 199 | 5 | 84 | 5 | 479 | 10 | 39173 | 257 | <LOD | 117 | <LOD | 22 | 1966 | 26 | 36418 | 220 | 69 | 17 | <LOD | 6 | 118 | 2 | 142 | 3 | 212 | 4 | 10.5 | 1.3 | 4.4 | 1.4 | 38.3 | 1.7 | 151 | 3 | 34 | 3 | 82 | 4 | <LOD | 29 | 183 | 52 | <LOD | 20 | <LOD | 24 | 7786 | 51 | <LOD | 30 | 17 | 2 | 23 | 3
OREAS 621 | <LOD | 359 | <LOD | 194 | <LOD | 105 | <LOD | 82 | 2778 | 33 | 126 | 8 | 13965 | 3183 | 24360 | 525 | 2022 | 177 | 23290 | 352 | 22805 | 307 | 3084 | 55 | 215 | 6 | 113 | 7 | 698 | 14 | 59738 | 437 | <LOD | 168 | <LOD | 29 | 4621 | 51 | 69000 | 474 | 184 | 26 | <LOD | 9 | 95 | 2 | 111 | 3 | 173 | 4 | 7.1 | 1.6 | 14.4 | 1.6 | 64 | 2 | 260 | 4 | 48 | 4 | 157 | 5 | 154 | 16 | <LOD | 243 | <LOD | 30 | <LOD | 37 | 13860 | 100 | <LOD | 40 | 25 | 2 | 15 | 4
OREAS 621 | <LOD | 361 | <LOD | 195 | <LOD | 106 | <LOD | 82 | 2771 | 33 | 142 | 8 | <LOD | 9173 | 24569 | 526 | 1747 | 171 | 23032 | 348 | 22333 | 301 | 2880 | 53 | 220 | 6 | 106 | 6 | 675 | 14 | 58989 | 425 | <LOD | 163 | <LOD | 30 | 4561 | 50 | 67853 | 459 | 184 | 25 | <LOD | 9 | 94 | 2 | 111 | 3 | 166 | 4 | 8.8 | 1.6 | 12.3 | 1.6 | 66 | 2 | 262 | 4 | 51 | 4 | 153 | 5 | 135 | 15 | <LOD | 237 | <LOD | 29 | <LOD | 36 | 13667 | 97 | <LOD | 39 | 21 | 2 | 29 | 4
OREAS 621 | <LOD | 350 | <LOD | 188 | <LOD | 102 | <LOD | 79 | 2804 | 32 | 157 | 8 | 16026 | 3130 | 24145 | 508 | 2070 | 173 | 23081 | 340 | 22564 | 297 | 2957 | 53 | 217 | 6 | 100 | 6 | 669 | 13 | 60489 | 441 | <LOD | 168 | <LOD | 31 | 4714 | 51 | 69731 | 477 | 206 | 26 | <LOD | 9 | 95 | 2 | 113 | 3 | 163 | 4 | 12.6 | 1.6 | 16.1 | 1.6 | 68 | 2 | 271 | 4 | 48 | 4 | 162 | 4 | 136 | 16 | <LOD | 244 | <LOD | 30 | <LOD | 37 | 13944 | 100 | <LOD | 40 | 25 | 2 | 26 | 4
OREAS 621 | <LOD | 360 | <LOD | 194 | <LOD | 104 | <LOD | 81 | 2721 | 33 | 147 | 8 | 18677 | 3206 | 23759 | 509 | 2294 | 179 | 22921 | 343 | 22181 | 296 | 3054 | 54 | 209 | 6 | 111 | 6 | 675 | 14 | 60060 | 434 | <LOD | 166 | <LOD | 29 | 4866 | 51 | 69078 | 469 | 147 | 25 | <LOD | 9 | 92 | 2 | 115 | 3 | 164 | 4 | 5 | 1.6 | 11.7 | 1.6 | 65 | 2 | 264 | 4 | 51 | 4 | 160 | 5 | <LOD | 45 | <LOD | 240 | <LOD | 30 | <LOD | 37 | 13934 | 99 | <LOD | 40 | 22 | 2 | 28 | 4
OREAS 621 | <LOD | 355 | <LOD | 191 | <LOD | 104 | <LOD | 81 | 2731 | 33 | 140 | 8 | 15770 | 3189 | 23559 | 512 | 2111 | 178 | 24022 | 359 | 22345 | 301 | 3084 | 55 | 226 | 6 | 109 | 6 | 678 | 14 | 62013 | 526 | <LOD | 198 | <LOD | 35 | 4716 | 60 | 71068 | 568 | 239 | 30 | <LOD | 11 | 94 | 3 | 117 | 3 | 170 | 4 | 11.4 | 1.9 | 11.8 | 1.6 | 68 | 2 | 260 | 4 | 45 | 4 | 160 | 5 | 209 | 18 | <LOD | 284 | <LOD | 35 | <LOD | 43 | 14153 | 119 | <LOD | 46 | 21 | 2 | 33 | 4
OREAS 621 | <LOD | 364 | <LOD | 195 | <LOD | 106 | <LOD | 82 | 2718 | 33 | 150 | 8 | 13269 | 3087 | 24292 | 514 | 2136 | 175 | 22825 | 340 | 21914 | 292 | 2943 | 53 | 216 | 6 | 104 | 6 | 641 | 13 | 59787 | 436 | <LOD | 167 | <LOD | 30 | 4883 | 52 | 69036 | 473 | 185 | 26 | <LOD | 9 | 92 | 2 | 115 | 3 | 167 | 4 | 6.6 | 1.6 | 13.7 | 1.6 | 64 | 2 | 274 | 4 | 52 | 4 | 172 | 5 | <LOD | 46 | <LOD | 242 | <LOD | 30 | <LOD | 37 | 13893 | 100 | <LOD | 40 | 20 | 2 | 28 | 4
OREAS 621 | <LOD | 353 | <LOD | 190 | <LOD | 103 | <LOD | 80 | 2777 | 33 | 129 | 8 | 19899 | 3306 | 25199 | 534 | 2036 | 178 | 23476 | 353 | 22872 | 307 | 3033 | 55 | 216 | 6 | 109 | 6 | 648 | 13 | 59745 | 437 | <LOD | 167 | <LOD | 30 | 4628 | 51 | 68753 | 472 | 191 | 26 | <LOD | 9 | 93 | 2 | 112 | 3 | 167 | 4 | 8 | 1.6 | 13.6 | 1.6 | 66 | 2 | 260 | 4 | 49 | 4 | 153 | 4 | 144 | 16 | <LOD | 242 | <LOD | 30 | <LOD | 37 | 13776 | 99 | <LOD | 40 | 22 | 2 | 25 | 4
OREAS 621 | <LOD | 351 | <LOD | 189 | <LOD | 103 | <LOD | 80 | 2791 | 33 | 144 | 8 | 16418 | 3211 | 24324 | 520 | 2279 | 181 | 23291 | 349 | 23090 | 308 | 2947 | 53 | 221 | 6 | 111 | 6 | 690 | 14 | 60398 | 427 | <LOD | 163 | <LOD | 30 | 4661 | 50 | 70042 | 465 | 167 | 25 | <LOD | 9 | 94 | 2 | 113 | 3 | 168 | 4 | 8.3 | 1.5 | 14.4 | 1.6 | 65 | 2 | 262 | 4 | 46 | 4 | 151 | 4 | 160 | 15 | <LOD | 237 | <LOD | 30 | <LOD | 36 | 13985 | 97 | <LOD | 39 | 20 | 2 | 22 | 4
oreas 621 | <LOD | 365 | <LOD | 196 | <LOD | 106 | <LOD | 83 | 2802 | 34 | 155 | 8 | 16456 | 3220 | 23761 | 519 | 2028 | 177 | 23170 | 352 | 22188 | 302 | 2882 | 53 | 213 | 6 | 104 | 6 | 662 | 14 | 60717 | 441 | <LOD | 167 | <LOD | 30 | 4762 | 51 | 69647 | 475 | 166 | 26 | <LOD | 9 | 91 | 2 | 111 | 3 | 168 | 4 | 10.6 | 1.6 | 12.3 | 1.6 | 64 | 2 | 261 | 4 | 47 | 4 | 159 | 5 | 86 | 15 | <LOD | 243 | <LOD | 30 | <LOD | 37 | 14030 | 100 | <LOD | 40 | 23 | 2 | 30 | 4
oreas 621 | <LOD | 363 | <LOD | 195 | <LOD | 106 | <LOD | 82 | 2741 | 33 | 165 | 8 | 13362 | 3077 | 23397 | 502 | 1841 | 169 | 22490 | 337 | 22045 | 293 | 2927 | 52 | 212 | 6 | 106 | 6 | 652 | 13 | 59924 | 437 | <LOD | 166 | <LOD | 30 | 4953 | 52 | 68887 | 472 | 170 | 26 | <LOD | 9 | 91 | 2 | 118 | 4 | 166 | 4 | 7.1 | 1.6 | 14.2 | 1.6 | 69 | 2 | 262 | 4 | 52 | 4 | 164 | 5 | <LOD | 45 | <LOD | 243 | <LOD | 30 | <LOD | 37 | 13762 | 99 | <LOD | 40 | 17 | 2 | 27 | 4
Oreas 621 | <LOD | 365 | <LOD | 197 | <LOD | 107 | <LOD | 83 | 2729 | 33 | 142 | 8 | 25758 | 3759 | 26856 | 592 | 2062 | 196 | 26459 | 413 | 26240 | 367 | 3399 | 63 | 256 | 7 | 121 | 7 | 756 | 16 | 60021 | 440 | <LOD | 169 | <LOD | 30 | 4632 | 51 | 69090 | 476 | 164 | 26 | <LOD | 9 | 99 | 2 | 108 | 3 | 176 | 4 | 7.9 | 1.6 | 13.9 | 1.7 | 64 | 2 | 260 | 4 | 50 | 4 | 166 | 5 | 151 | 16 | <LOD | 244 | <LOD | 30 | <LOD | 37 | 13827 | 100 | <LOD | 40 | 20 | 2 | 28 | 4
Oreas 621 | <LOD | 362 | <LOD | 195 | <LOD | 106 | <LOD | 82 | 2750 | 33 | 144 | 8 | 15016 | 3208 | 22792 | 510 | 2070 | 180 | 23290 | 356 | 22146 | 304 | 2987 | 55 | 222 | 6 | 117 | 7 | 675 | 14 | 58622 | 431 | <LOD | 167 | <LOD | 30 | 4544 | 51 | 67705 | 467 | 160 | 26 | <LOD | 9 | 92 | 2 | 114 | 3 | 164 | 4 | 10.2 | 1.6 | 11.6 | 1.6 | 65 | 2 | 259 | 4 | 51 | 4 | 158 | 5 | 126 | 15 | <LOD | 242 | <LOD | 30 | <LOD | 37 | 13608 | 99 | <LOD | 40 | 23 | 2 | 26 | 4
OREAS 621 | <LOD | 358 | <LOD | 193 | <LOD | 105 | <LOD | 81 | 2724 | 33 | 155 | 8 | 15472 | 3146 | 22948 | 501 | 2341 | 180 | 23103 | 346 | 22330 | 299 | 2970 | 53 | 221 | 6 | 109 | 6 | 657 | 13 | 59684 | 435 | <LOD | 167 | <LOD | 30 | 4637 | 51 | 68356 | 468 | 193 | 25 | <LOD | 9 | 92 | 2 | 113 | 3 | 161 | 4 | 7.7 | 1.6 | 14.5 | 1.6 | 67 | 2 | 266 | 4 | 53 | 4 | 164 | 5 | 91 | 15 | <LOD | 242 | <LOD | 30 | <LOD | 37 | 13697 | 98 | <LOD | 40 | 17 | 2 | 31 | 4
OREAS 622 | <LOD | 428 | <LOD | 230 | <LOD | 125 | <LOD | 97 | 2460 | 36 | 226 | 11 | 29935 | 4358 | 45163 | 877 | 3460 | 246 | 21069 | 369 | 26737 | 392 | 3376 | 66 | 233 | 7 | 140 | 8 | 824 | 18 | 89690 | 782 | <LOD | 252 | <LOD | 48 | 7064 | 89 | 157662 | 1311 | 401 | 40 | <LOD | 15 | 75 | 3 | 84 | 4 | 131 | 4 | 15 | 2 | 21 | 2 | 93 | 3 | 422 | 6 | 100 | 5 | 227 | 6 | 349 | 28 | <LOD | 441 | <LOD | 48 | 122 | 22 | 22415 | 193 | <LOD | 57 | 29 | 2 | 49 | 6
OREAS 622 | <LOD | 425 | <LOD | 229 | <LOD | 124 | <LOD | 96 | 2557 | 37 | 208 | 11 | 33467 | 4570 | 46214 | 914 | 3871 | 262 | 21159 | 380 | 27294 | 410 | 3478 | 70 | 223 | 7 | 131 | 8 | 832 | 18 | 88990 | 774 | <LOD | 250 | <LOD | 48 | 7214 | 89 | 157797 | 1308 | 352 | 39 | 16 | 5 | 73 | 3 | 89 | 4 | 128 | 4 | 13 | 2 | 22 | 2 | 96 | 3 | 434 | 7 | 94 | 5 | 218 | 6 | 303 | 28 | <LOD | 439 | <LOD | 47 | 119 | 22 | 22302 | 191 | <LOD | 56 | 32 | 3 | 45 | 6
OREAS 622 | <LOD | 443 | <LOD | 239 | <LOD | 131 | <LOD | 101 | 2427 | 38 | 202 | 12 | 32677 | 4691 | 44902 | 930 | 3655 | 267 | 20738 | 390 | 26693 | 418 | 3236 | 69 | 232 | 8 | 142 | 8 | 824 | 19 | 85199 | 754 | <LOD | 247 | <LOD | 46 | 6844 | 87 | 150994 | 1272 | 293 | 39 | <LOD | 15 | 68 | 3 | 94 | 4 | 122 | 4 | 13 | 2 | 18 | 2 | 98 | 3 | 421 | 7 | 101 | 5 | 226 | 6 | 294 | 28 | <LOD | 435 | <LOD | 47 | 106 | 22 | 21555 | 188 | <LOD | 56 | 31 | 3 | 63 | 6
OREAS 622 | <LOD | 430 | <LOD | 230 | <LOD | 125 | <LOD | 96 | 2571 | 37 | 224 | 11 | 30742 | 4474 | 45969 | 902 | 4369 | 268 | 21533 | 381 | 27564 | 409 | 3335 | 67 | 240 | 8 | 141 | 8 | 859 | 18 | 89484 | 815 | <LOD | 263 | <LOD | 50 | 7321 | 94 | 158387 | 1375 | 457 | 41 | <LOD | 16 | 69 | 3 | 93 | 4 | 128 | 4 | 12 | 2 | 21 | 2 | 104 | 3 | 425 | 6 | 103 | 5 | 225 | 6 | 280 | 29 | <LOD | 462 | <LOD | 49 | 105 | 23 | 22209 | 200 | <LOD | 59 | 31 | 3 | 46 | 6
OREAS 622 | <LOD | 426 | <LOD | 228 | <LOD | 124 | <LOD | 95 | 2411 | 36 | 215 | 11 | 29767 | 4505 | 47012 | 922 | 3473 | 254 | 21498 | 384 | 27707 | 414 | 3486 | 69 | 221 | 7 | 150 | 8 | 843 | 18 | 89576 | 777 | <LOD | 249 | <LOD | 47 | 7671 | 91 | 157674 | 1305 | 402 | 39 | <LOD | 15 | 72 | 3 | 81 | 4 | 123 | 4 | 13 | 2 | 23 | 2 | 95 | 3 | 412 | 6 | 101 | 5 | 230 | 6 | <LOD | 82 | <LOD | 439 | <LOD | 47 | <LOD | 66 | 22287 | 191 | <LOD | 57 | 36 | 3 | 48 | 6
OREAS 622 | <LOD | 430 | <LOD | 230 | <LOD | 126 | <LOD | 97 | 2569 | 37 | 217 | 11 | 34384 | 4554 | 46307 | 911 | 3687 | 257 | 20943 | 376 | 26807 | 402 | 3491 | 69 | 220 | 7 | 139 | 8 | 828 | 18 | 88120 | 765 | <LOD | 249 | <LOD | 46 | 7062 | 88 | 155627 | 1288 | 358 | 39 | <LOD | 15 | 72 | 3 | 81 | 4 | 131 | 4 | 16 | 2 | 21 | 2 | 99 | 3 | 428 | 7 | 100 | 5 | 226 | 6 | 327 | 28 | <LOD | 435 | <LOD | 47 | 137 | 22 | 22164 | 190 | <LOD | 56 | 37 | 3 | 48 | 6
Oreas 622 | <LOD | 436 | <LOD | 234 | <LOD | 127 | <LOD | 98 | 2431 | 37 | 219 | 12 | 32798 | 5221 | 53425 | 1108 | 4167 | 303 | 24397 | 464 | 31634 | 510 | 3906 | 83 | 267 | 9 | 166 | 10 | 1021 | 23 | 87538 | 764 | <LOD | 248 | <LOD | 47 | 6890 | 88 | 154820 | 1288 | 377 | 39 | <LOD | 15 | 73 | 3 | 84 | 4 | 124 | 4 | 14 | 2 | 19 | 2 | 95 | 3 | 414 | 7 | 86 | 5 | 219 | 6 | 383 | 28 | <LOD | 436 | <LOD | 47 | 131 | 22 | 21814 | 188 | <LOD | 56 | 34 | 3 | 48 | 6
Oreas 622 | <LOD | 432 | <LOD | 230 | <LOD | 126 | <LOD | 97 | 2530 | 37 | 195 | 11 | 23757 | 4438 | 47134 | 933 | 3985 | 267 | 21605 | 389 | 27549 | 416 | 3448 | 70 | 230 | 8 | 153 | 8 | 831 | 18 | 87576 | 747 | <LOD | 245 | <LOD | 45 | 7148 | 87 | 156131 | 1268 | 307 | 38 | <LOD | 15 | 70 | 3 | 88 | 4 | 123 | 4 | 11.8 | 2 | 22 | 2 | 97 | 3 | 423 | 7 | 101 | 5 | 227 | 6 | 291 | 27 | <LOD | 428 | <LOD | 46 | 102 | 22 | 22179 | 186 | <LOD | 55 | 37 | 3 | 43 | 6
OREAS 622 | <LOD | 437 | <LOD | 235 | <LOD | 128 | <LOD | 98 | 2494 | 38 | 210 | 12 | 30786 | 4473 | 44859 | 893 | 3762 | 258 | 20189 | 367 | 27072 | 406 | 3326 | 67 | 239 | 8 | 140 | 8 | 794 | 18 | 88084 | 759 | <LOD | 246 | <LOD | 47 | 7471 | 89 | 155113 | 1274 | 385 | 39 | <LOD | 15 | 68 | 3 | 82 | 4 | 129 | 4 | 9.5 | 2 | 24 | 2 | 96 | 3 | 415 | 7 | 100 | 5 | 228 | 6 | 107 | 27 | <LOD | 432 | <LOD | 46 | <LOD | 65 | 21947 | 187 | <LOD | 56 | 35 | 3 | 53 | 6
OREAS 623 | <LOD | 345 | <LOD | 186 | <LOD | 104 | <LOD | 81 | 1383 | 23 | 53 | 4 | 17303 | 4050 | 44711 | 842 | 606 | 198 | 11535 | 254 | 14928 | 242 | 1974 | 56 | 131 | 6 | 133 | 8 | 677 | 16 | 197129 | 1284 | <LOD | 280 | <LOD | 24 | 15242 | 109 | 9440 | 69 | 78 | 10 | 19.8 | 1.6 | 59.8 | 1.6 | 81 | 3 | 137 | 3 | 12.5 | 1.4 | 9.5 | 1.3 | 22.5 | 1.7 | 47 | 2 | 31 | 4 | 28 | 4 | 17 | 6 | <LOD | 86 | <LOD | 14 | 17 | 5 | 2188 | 19 | <LOD | 26 | 14.1 | 2 | <LOD | 8
oreas 623 | <LOD | 348 | <LOD | 187 | 113 | 35 | <LOD | 81 | 1429 | 24 | 39 | 4 | <LOD | 11426 | 44106 | 827 | <LOD | 579 | 11746 | 253 | 14932 | 240 | 1905 | 54 | 126 | 6 | 137 | 8 | 660 | 16 | 197783 | 1298 | <LOD | 281 | <LOD | 25 | 15324 | 110 | 9491 | 70 | 83 | 10 | 20.3 | 1.6 | 59.3 | 1.6 | 80 | 3 | 134 | 3 | 12.7 | 1.4 | 8.2 | 1.3 | 22 | 1.6 | 49 | 2 | 31 | 4 | 34 | 4 | <LOD | 16 | <LOD | 86 | <LOD | 14 | 27 | 5 | 2182 | 20 | <LOD | 27 | 13.9 | 2 | 9 | 3
OREAS 623 | <LOD | 343 | <LOD | 185 | <LOD | 104 | <LOD | 81 | 1425 | 24 | 44 | 4 | 13639 | 4004 | 44256 | 840 | 1240 | 211 | 11303 | 252 | 15222 | 247 | 2007 | 56 | 132 | 6 | 122 | 8 | 684 | 16 | 198729 | 1303 | <LOD | 283 | <LOD | 25 | 15552 | 112 | 9554 | 70 | 73 | 10 | 23.3 | 1.7 | 60.6 | 1.6 | 80 | 3 | 130 | 3 | 13.1 | 1.4 | 8.3 | 1.3 | 21.8 | 1.7 | 54 | 2 | 32 | 4 | 34 | 4 | <LOD | 17 | <LOD | 87 | <LOD | 15 | 17 | 5 | 2218 | 20 | <LOD | 27 | 15 | 2 | <LOD | 9
OREAS 623 | <LOD | 354 | <LOD | 191 | <LOD | 107 | <LOD | 83 | 1361 | 24 | 47 | 4 | <LOD | 13623 | 43159 | 984 | <LOD | 730 | 11318 | 303 | 14982 | 290 | 2030 | 67 | 109 | 7 | 121 | 10 | 629 | 19 | 200181 | 1329 | <LOD | 287 | <LOD | 25 | 15530 | 113 | 9636 | 72 | 74 | 10 | 20.8 | 1.7 | 60.1 | 1.7 | 84 | 3 | 143 | 3 | 14.8 | 1.5 | 10.3 | 1.4 | 25.6 | 1.7 | 51 | 2 | 34 | 4 | 32 | 4 | <LOD | 17 | <LOD | 88 | <LOD | 15 | 27 | 5 | 2236 | 20 | <LOD | 27 | 13 | 2 | 9 | 3
OREAS 623 | <LOD | 343 | <LOD | 184 | <LOD | 103 | <LOD | 80 | 1359 | 23 | 45 | 4 | <LOD | 11209 | 43925 | 795 | 894 | 193 | 11575 | 242 | 15097 | 233 | 1984 | 53 | 110 | 6 | 131 | 8 | 677 | 15 | 200386 | 1308 | <LOD | 283 | <LOD | 25 | 15634 | 112 | 9643 | 71 | 98 | 10 | 19.9 | 1.6 | 60.1 | 1.6 | 80 | 3 | 142 | 3 | 16 | 1.5 | 7.2 | 1.3 | 24 | 1.7 | 48 | 2 | 32 | 4 | 33 | 4 | 23 | 6 | <LOD | 87 | <LOD | 15 | 26 | 5 | 2184 | 20 | <LOD | 26 | 12.7 | 1.9 | <LOD | 9
OREAS 623 | <LOD | 332 | <LOD | 179 | <LOD | 101 | <LOD | 78 | 1170 | 22 | 44 | 3 | 12835 | 3950 | 45038 | 843 | 742 | 199 | 11682 | 254 | 14859 | 240 | 1923 | 55 | 111 | 6 | 122 | 8 | 672 | 16 | 189951 | 1193 | <LOD | 264 | <LOD | 23 | 14887 | 103 | 9027 | 64 | 80 | 9 | 21 | 1.6 | 54.5 | 1.5 | 76 | 3 | 128 | 3 | 10.3 | 1.4 | 10.7 | 1.3 | 16.2 | 1.6 | 47 | 2 | 27 | 4 | 26 | 4 | <LOD | 15 | <LOD | 81 | <LOD | 14 | 13 | 4 | 2073 | 18 | <LOD | 25 | 13.4 | 1.9 | <LOD | 8
OREAS 623 | <LOD | 387 | <LOD | 207 | <LOD | 116 | <LOD | 89 | 1099 | 24 | 78 | 5 | <LOD | 14512 | 69156 | 1256 | <LOD | 744 | 4505 | 210 | 17984 | 304 | 1876 | 62 | 97 | 7 | 137 | 10 | 726 | 19 | 279526 | 2051 | <LOD | 375 | <LOD | 33 | 29703 | 228 | 23648 | 181 | 117 | 17 | 44 | 3 | 34.4 | 1.7 | 42 | 2 | 92 | 3 | 11.3 | 1.7 | 14.6 | 1.4 | 42 | 2 | 112 | 3 | 42 | 4 | 73 | 4 | <LOD | 28 | <LOD | 149 | <LOD | 23 | 71 | 8 | 5327 | 45 | <LOD | 34 | 16 | 2 | 11 | 3
OREAS 623 | <LOD | 357 | <LOD | 192 | <LOD | 107 | <LOD | 83 | 1392 | 24 | 42 | 4 | 14492 | 3948 | 44076 | 828 | 610 | 196 | 11218 | 248 | 14454 | 235 | 1953 | 55 | 114 | 6 | 133 | 8 | 642 | 16 | 200023 | 1320 | <LOD | 284 | <LOD | 25 | 15515 | 112 | 9584 | 71 | 85 | 10 | 25.7 | 1.7 | 59.6 | 1.7 | 85 | 3 | 145 | 3 | 12.7 | 1.5 | 10.8 | 1.4 | 22.2 | 1.7 | 48 | 2 | 31 | 4 | 33 | 4 | <LOD | 17 | <LOD | 87 | <LOD | 15 | 25 | 5 | 2168 | 20 | <LOD | 27 | 13 | 2 | 11 | 3
OREAS 623 | <LOD | 344 | <LOD | 186 | <LOD | 105 | <LOD | 81 | 1379 | 23 | 42 | 4 | <LOD | 11319 | 41380 | 785 | 1164 | 201 | 10684 | 239 | 14519 | 233 | 1934 | 54 | 121 | 6 | 125 | 8 | 665 | 16 | 200592 | 1311 | <LOD | 283 | <LOD | 25 | 15622 | 112 | 9701 | 71 | 92 | 10 | 20.6 | 1.6 | 58.5 | 1.6 | 80 | 3 | 133 | 3 | 14.9 | 1.5 | 8.9 | 1.3 | 19.2 | 1.7 | 50 | 2 | 28 | 4 | 26 | 4 | 18 | 6 | <LOD | 87 | <LOD | 15 | 25 | 5 | 2208 | 20 | <LOD | 27 | 14 | 2 | 13 | 3
Oreas 623 | <LOD | 355 | <LOD | 192 | <LOD | 108 | <LOD | 83 | 1460 | 24 | 47 | 4 | 19360 | 4659 | 53186 | 1024 | 1734 | 248 | 13512 | 302 | 18078 | 302 | 2335 | 66 | 143 | 7 | 140 | 9 | 779 | 19 | 206706 | 1361 | <LOD | 290 | <LOD | 26 | 16027 | 115 | 9866 | 73 | 82 | 10 | 23.2 | 1.7 | 60.6 | 1.7 | 85 | 3 | 145 | 3 | 14.6 | 1.5 | 10.7 | 1.4 | 19.9 | 1.7 | 49 | 2 | 34 | 4 | 42 | 4 | <LOD | 17 | <LOD | 88 | <LOD | 15 | 27 | 5 | 2273 | 20 | <LOD | 27 | 13 | 2 | 13 | 3
Oreas 623 | <LOD | 343 | <LOD | 186 | 137 | 35 | <LOD | 81 | 1380 | 23 | 45 | 4 | <LOD | 11316 | 43297 | 810 | 657 | 194 | 11144 | 244 | 14771 | 236 | 1896 | 54 | 120 | 6 | 122 | 8 | 656 | 16 | 198220 | 1286 | <LOD | 280 | <LOD | 25 | 15412 | 109 | 9522 | 69 | 102 | 10 | 19.8 | 1.6 | 60.7 | 1.6 | 80 | 3 | 139 | 3 | 12.1 | 1.4 | 10.2 | 1.3 | 19.7 | 1.7 | 48 | 2 | 36 | 4 | 30 | 4 | 22 | 6 | <LOD | 86 | <LOD | 15 | 19 | 5 | 2181 | 19 | <LOD | 26 | 13.5 | 2 | 12 | 3
OREAS 623 | <LOD | 351 | <LOD | 189 | <LOD | 106 | <LOD | 82 | 1394 | 24 | 44 | 4 | <LOD | 11285 | 42822 | 801 | 1055 | 200 | 11359 | 246 | 14613 | 234 | 1839 | 53 | 127 | 6 | 127 | 8 | 656 | 16 | 204516 | 1565 | <LOD | 331 | <LOD | 29 | 15962 | 133 | 9784 | 84 | 70 | 12 | 22.4 | 2 | 64.3 | 1.9 | 80 | 3 | 140 | 3 | 12.8 | 1.7 | 8.5 | 1.3 | 21.5 | 1.7 | 50 | 2 | 32 | 4 | 32 | 4 | <LOD | 19 | <LOD | 102 | <LOD | 18 | 21 | 5 | 2254 | 23 | <LOD | 31 | 16 | 2 | 10 | 3
OREAS 624 | <LOD | 394 | <LOD | 210 | <LOD | 117 | <LOD | 90 | 1149 | 24 | 72 | 5 | 20134 | 4724 | 64900 | 1158 | 1485 | 244 | 6159 | 207 | 16920 | 280 | 1760 | 59 | 114 | 7 | 144 | 9 | 721 | 18 | 285178 | 2066 | <LOD | 374 | <LOD | 33 | 30181 | 229 | 24043 | 182 | 157 | 17 | 42 | 3 | 33.2 | 1.7 | 39 | 2 | 94 | 3 | 13.7 | 1.6 | 17.4 | 1.5 | 39 | 2 | 116 | 3 | 52 | 4 | 76 | 4 | <LOD | 28 | <LOD | 148 | <LOD | 22 | 69 | 8 | 5393 | 45 | <LOD | 34 | 14 | 2 | 14 | 4
oreas 624 | <LOD | 404 | <LOD | 215 | <LOD | 120 | <LOD | 92 | 1181 | 25 | 90 | 6 | 17652 | 4715 | 65453 | 1170 | 1400 | 244 | 6135 | 207 | 17128 | 284 | 1719 | 58 | 96 | 7 | 148 | 9 | 706 | 18 | 286852 | 2109 | <LOD | 380 | <LOD | 35 | 30478 | 234 | 24102 | 185 | 149 | 17 | 42 | 3 | 33.3 | 1.7 | 46 | 2 | 95 | 3 | 10.7 | 1.7 | 15.2 | 1.5 | 45 | 2 | 117 | 3 | 56 | 4 | 79 | 4 | <LOD | 28 | <LOD | 150 | <LOD | 23 | 66 | 8 | 5443 | 46 | <LOD | 34 | 14 | 2 | 14 | 4
OREAS 624 | <LOD | 396 | <LOD | 212 | <LOD | 119 | <LOD | 92 | 1119 | 24 | 76 | 5 | 18625 | 4998 | 67125 | 1252 | 1585 | 261 | 6815 | 227 | 17843 | 308 | 1877 | 63 | 114 | 7 | 170 | 10 | 768 | 20 | 278156 | 2062 | <LOD | 378 | <LOD | 33 | 29383 | 228 | 23469 | 182 | 147 | 17 | 49 | 3 | 33.5 | 1.7 | 41 | 2 | 95 | 3 | 11.7 | 1.7 | 18.3 | 1.5 | 41 | 2 | 114 | 3 | 48 | 4 | 71 | 4 | <LOD | 28 | <LOD | 148 | <LOD | 23 | 89 | 8 | 5277 | 45 | <LOD | 34 | 12 | 2 | <LOD | 10
OREAS 624 | <LOD | 396 | <LOD | 211 | <LOD | 118 | <LOD | 91 | 1176 | 24 | 82 | 6 | 16816 | 4799 | 65725 | 1194 | 1832 | 257 | 6186 | 212 | 17353 | 291 | 1845 | 61 | 91 | 7 | 167 | 10 | 725 | 18 | 281622 | 2066 | <LOD | 377 | <LOD | 33 | 29717 | 228 | 23940 | 183 | 143 | 17 | 38 | 2 | 32.1 | 1.7 | 41 | 2 | 92 | 3 | 14.6 | 1.7 | 19.3 | 1.5 | 45 | 2 | 117 | 3 | 51 | 4 | 72 | 4 | <LOD | 29 | <LOD | 149 | <LOD | 22 | 76 | 8 | 5364 | 45 | <LOD | 34 | 16 | 2 | <LOD | 10
OREAS 624 | <LOD | 398 | <LOD | 213 | <LOD | 119 | <LOD | 91 | 1146 | 25 | 73 | 5 | 24623 | 4820 | 64197 | 1156 | 1606 | 247 | 6162 | 208 | 17018 | 283 | 1792 | 59 | 95 | 7 | 155 | 9 | 727 | 18 | 277272 | 2014 | <LOD | 371 | <LOD | 32 | 29464 | 224 | 23434 | 178 | 128 | 17 | 46 | 3 | 34 | 1.7 | 47 | 2 | 93 | 3 | 11.1 | 1.6 | 15.7 | 1.5 | 44 | 2 | 118 | 3 | 51 | 4 | 75 | 4 | <LOD | 28 | <LOD | 145 | <LOD | 22 | 94 | 8 | 5275 | 44 | <LOD | 34 | 14 | 2 | <LOD | 11
OREAS 624 | <LOD | 395 | <LOD | 210 | <LOD | 118 | <LOD | 90 | 1129 | 24 | 71 | 5 | 24232 | 4930 | 67140 | 1209 | 1735 | 255 | 6362 | 214 | 17317 | 290 | 1832 | 61 | 96 | 7 | 151 | 9 | 742 | 19 | 282632 | 2069 | <LOD | 378 | <LOD | 33 | 30069 | 230 | 24044 | 183 | 135 | 17 | 45 | 3 | 31.9 | 1.7 | 39 | 2 | 90 | 3 | 9 | 1.6 | 19.3 | 1.6 | 42 | 2 | 115 | 3 | 44 | 4 | 74 | 4 | <LOD | 28 | <LOD | 149 | <LOD | 22 | 81 | 8 | 5401 | 46 | <LOD | 34 | 19 | 2 | 11 | 4
OREAS 624 | <LOD | 388 | <LOD | 208 | <LOD | 116 | <LOD | 89 | 1109 | 24 | 86 | 5 | 15076 | 4718 | 65448 | 1179 | 1742 | 252 | 6140 | 209 | 17402 | 289 | 1781 | 60 | 95 | 7 | 150 | 9 | 748 | 18 | 285847 | 2098 | <LOD | 380 | <LOD | 33 | 30344 | 233 | 24088 | 184 | 136 | 17 | 45 | 3 | 33.4 | 1.7 | 43 | 2 | 94 | 3 | 13.2 | 1.7 | 17.6 | 1.5 | 36 | 2 | 115 | 3 | 45 | 4 | 72 | 4 | <LOD | 29 | <LOD | 150 | <LOD | 23 | 83 | 8 | 5424 | 46 | <LOD | 34 | 12.7 | 2 | <LOD | 10
Oreas 624 | <LOD | 390 | <LOD | 208 | <LOD | 117 | <LOD | 90 | 1151 | 24 | 79 | 6 | 24444 | 5839 | 80494 | 1558 | 2219 | 310 | 7369 | 260 | 21018 | 381 | 2214 | 75 | 118 | 8 | 189 | 11 | 926 | 24 | 293379 | 2194 | <LOD | 392 | <LOD | 34 | 31190 | 243 | 24644 | 192 | 147 | 18 | 40 | 3 | 34.7 | 1.7 | 46 | 2 | 99 | 3 | 10.5 | 1.7 | 18.2 | 1.5 | 43 | 2 | 121 | 3 | 41 | 4 | 80 | 4 | <LOD | 30 | <LOD | 154 | <LOD | 23 | 88 | 8 | 5554 | 48 | <LOD | 35 | 15 | 2 | 11 | 3
Oreas 624 | <LOD | 392 | <LOD | 210 | <LOD | 117 | <LOD | 90 | 1139 | 24 | 81 | 5 | 15734 | 4699 | 65275 | 1171 | 1205 | 241 | 6038 | 207 | 17094 | 284 | 1772 | 59 | 107 | 7 | 155 | 9 | 718 | 18 | 282012 | 2027 | <LOD | 370 | <LOD | 33 | 29771 | 224 | 23708 | 178 | 136 | 17 | 46 | 3 | 29.3 | 1.6 | 41 | 2 | 97 | 3 | 14.4 | 1.6 | 15.7 | 1.5 | 43 | 2 | 116 | 3 | 45 | 4 | 73 | 4 | <LOD | 28 | <LOD | 145 | <LOD | 22 | 88 | 8 | 5347 | 44 | <LOD | 33 | 18 | 2 | <LOD | 10
OREAS 624 | <LOD | 399 | <LOD | 213 | <LOD | 119 | <LOD | 91 | 1192 | 25 | 77 | 6 | 17274 | 4777 | 67266 | 1203 | 1587 | 250 | 6068 | 209 | 17245 | 287 | 1820 | 60 | 93 | 7 | 162 | 9 | 720 | 18 | 276780 | 1999 | <LOD | 368 | <LOD | 33 | 29290 | 221 | 23425 | 177 | 157 | 17 | 40 | 2 | 32.8 | 1.7 | 44 | 2 | 98 | 3 | 7.9 | 1.6 | 18.6 | 1.5 | 43 | 2 | 117 | 3 | 58 | 4 | 72 | 4 | <LOD | 27 | <LOD | 146 | <LOD | 22 | 62 | 8 | 5228 | 44 | <LOD | 33 | 14 | 2 | <LOD | 10
Medianvalues
| Nd | Nd | Pr | Pr | Ce | Ce | La | La | Ba | Ba | Y | Y | P | P | S | S | Cl | Cl | K | K | Ca | Ca | Ti | Ti | V | V | Cr | Cr | Mn | Mn | Fe | Fe | Co | Co | Ni | Ni | Cu | Cu | Zn | Zn | As | As | Se | Se | Rb | Rb | Sr | Sr | Zr | Zr | Nb | Nb | Mo | Mo | Ag | Ag | Cd | Cd | Sn | Sn | Sb | Sb | Ta | Ta | W | W | Au | Au | Hg | Hg | Pb | Pb | Bi | Bi | Th | Th | U | U
OREAS 620 | #NUM! | 312 | #NUM! | 169 | 117.5 | 92 | #NUM! | 72 | 2801 | 30 | 93 | 5 | 9471 | 2418 | 12689 | 323 | 1170 | 131 | 25117 | 333 | 16784 | 218 | 2556 | 43 | 201 | 5 | 85 | 5 | 498 | 10 | 39377 | 258 | #NUM! | 118 | #NUM! | 21 | 1905 | 27 | 36550 | 220 | 97 | 17 | #NUM! | 6 | 118 | 2 | 148 | 3 | 208 | 4 | 11.3 | 1.4 | 6.7 | 1.4 | 38.5 | 1.7 | 153 | 3 | 39 | 3 | 91 | 4 | 89 | 10 | 200 | 154 | #NUM! | 20 | #NUM! | 24 | 7804 | 51 | #NUM! | 30 | 17 | 2 | 20 | 3
OREAS 621 | #NUM! | 360 | #NUM! | 194 | #NUM! | 105 | #NUM! | 82 | 2750 | 33 | 144 | 8 | 15898 | 3206 | 24145 | 514 | 2070 | 178 | 23170 | 349 | 22333 | 301 | 2970 | 53 | 217 | 6 | 109 | 6 | 675 | 14 | 59924 | 437 | #NUM! | 167 | #NUM! | 30 | 4661 | 51 | 69036 | 472 | 184 | 26 | #NUM! | 9 | 93 | 2 | 113 | 3 | 167 | 4 | 8 | 1.6 | 13.7 | 1.6 | 65 | 2 | 262 | 4 | 50 | 4 | 160 | 5 | 140 | 16 | #NUM! | 242 | #NUM! | 30 | #NUM! | 37 | 13860 | 99 | #NUM! | 40 | 21 | 2 | 28 | 4
OREAS 622 | #NUM! | 430 | #NUM! | 230 | #NUM! | 126 | #NUM! | 97 | 2494 | 37 | 215 | 11 | 30786 | 4505 | 46214 | 914 | 3762 | 262 | 21159 | 381 | 27294 | 410 | 3448 | 69 | 232 | 8 | 141 | 8 | 831 | 18 | 88120 | 765 | #NUM! | 249 | #NUM! | 47 | 7148 | 89 | 156131 | 1288 | 377 | 39 | 16 | 15 | 72 | 3 | 84 | 4 | 128 | 4 | 13 | 2 | 21 | 2 | 96 | 3 | 422 | 7 | 100 | 5 | 226 | 6 | 298.5 | 28 | #NUM! | 436 | #NUM! | 47 | 119 | 22 | 22179 | 190 | #NUM! | 56 | 34 | 3 | 48 | 6
OREAS 623 | #NUM! | 346.5 | #NUM! | 186.5 | 125 | 104.5 | #NUM! | 81 | 1381.5 | 24 | 44.5 | 4 | 14492 | 11247 | 44091 | 834 | 894 | 200.5 | 11338.5 | 250 | 14930 | 240 | 1943.5 | 55 | 120.5 | 6 | 129 | 8 | 668.5 | 16 | 200102 | 1309.5 | #NUM! | 283 | #NUM! | 25 | 15541 | 112 | 9610 | 71 | 82.5 | 10 | 20.9 | 1.65 | 59.95 | 1.6 | 80 | 3 | 138 | 3 | 12.75 | 1.5 | 9.85 | 1.3 | 21.9 | 1.7 | 49 | 2 | 32 | 4 | 32.5 | 4 | 20 | 16.5 | #NUM! | 87 | #NUM! | 15 | 25 | 5 | 2198 | 20 | #NUM! | 27 | 13.7 | 2 | 11 | 3
OREAS 624 | #NUM! | 395.5 | #NUM! | 210.5 | #NUM! | 118 | #NUM! | 90.5 | 1147.5 | 24 | 78 | 5 | 18138.5 | 4788 | 65589 | 1186.5 | 1596.5 | 251 | 6160.5 | 209 | 17281 | 288 | 1806 | 60 | 96 | 7 | 155 | 9 | 726 | 18 | 282322 | 2066 | #NUM! | 377.5 | #NUM! | 33 | 29920 | 228.5 | 23991.5 | 182.5 | 145 | 17 | 43.5 | 3 | 33.25 | 1.7 | 42 | 2 | 94.5 | 3 | 11.4 | 1.65 | 17.9 | 1.5 | 43 | 2 | 116.5 | 3 | 49.5 | 4 | 73.5 | 4 | #NUM! | 28 | #NUM! | 148.5 | #NUM! | 22 | 82 | 8 | 5378.5 | 45 | #NUM! | 34 | 14 | 2 | 12.5 | 10

## pXRF vs. standards
As | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 53.703 | 97
| 621 | 84.674 | 184
| 622 | 110.3137 | 377
| 623 | 86.0053 | 82.5
| 624 | 114.5497 | 145
Fe | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 3.0128 | 39377 | 3.9377
| 621 | 3.7134 | 59787 | 5.9787
| 622 | 4.2585 | 88120 | 8.812
| 623 | 13.447 | 200102 | 20.0102
| 624 | 16.3109 | 282322 | 28.2322
Zn | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 3.1408 | 36550 | 3.655
| 621 | 5.2174 | 69036 | 6.9036
| 622 | 9.8985 | 156131 | 15.6131
| 623 | 1.0348 | 9610 | 0.961
| 624 | 2.4129 | 23991.5 | 2.3992
S | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 2.49 | 12689 | 1.2689
| 621 | 4.506 | 24145 | 2.4145
| 622 | 7.6778 | 46214 | 4.6214
| 623 | 8.8333 | 44091 | 4.4091
| 624 | 13.18 | 65589 | 6.5589
Ca | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 1.6319 | 16784 | 1.6784
| 621 | 2.0041 | 22333 | 2.2333
| 622 | 2.1368 | 27294 | 2.7294
| 623 | 1.3699 | 14930 | 1.493
| 624 | 1.4886 | 17281 | 1.7281
Pb | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 0.77175 | 7804 | 0.7804
| 621 | 1.3254 | 13893 | 1.3893
| 622 | 2.1053 | 22179 | 2.2179
| 623 | 0.24727 | 2198 | 0.2198
| 624 | 0.61199 | 5378.5 | 0.53785
Ti | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 0.15499 | 2556 | 0.2556
| 621 | 0.18101 | 2947 | 0.2947
| 622 | 0.18723 | 3448 | 0.3448
| 623 | 0.15009 | 1943.5 | 0.19435
| 624 | 0.14567 | 1806 | 0.1806
Mn | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 0.044926 | 498 | 0.0498
| 621 | 0.055401 | 669 | 0.0669
| 622 | 0.061277 | 831 | 0.0831
| 623 | 0.060477 | 668.5 | 0.06685
| 624 | 0.065984 | 726 | 0.0726
Rb | OREAS no. | OREAS standard | pXRF median value | Converted to % wt.
| 620 | 122.534 | 118 | 0.0118
| 621 | 88.7017 | 94 | 0.0094
| 622 | 57.5778 | 72 | 0.0072
| 623 | 66.6862 | 59.95 | 0.005995
| 624 | 33.0439 | 33.25 | 0.003325