Supplementary material 1. Table of portable X-ray fluorescence analyser performance. Certified value for certified reference material (xCRM), value for reference material (xRM), average elemental concentration determined by pXRF (xav), standard deviation (σ), relative standanrd deviation (RSD), limit of detection (LOD), limit of quantification (LOQ) and percentage relative difference (RD) for all CRMs.

|  | BCS-CRM No.393 |  |  |  |  |  |  | EURONORM-CRM No.782-1 |  |  |  |  |  |  | NCS DC 14020a |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | xCRM | xav | σ | RSD | LOD | LOQ | RD | xCRM | xav | σ | RSD | LOD | LOQ | RD | xCRM | xav | σ | RSD | LOD | LOQ | RD |
| Mg (%) | 0.09 | 0.34 | 0.68 | 200.02 | 2.04 | 6.73 | 275.86 | 12.84 | 14.19 | 0.82 | 5.76 | 2.45 | 8.1 | 10.54 | 9.28 | 9.54 | 0.41 | 4.31 | 1.23 | 4.07 | 2.86 |
| S (ppm) | 700 | bd | - | - | - | - | - | 489 | 146 | 75.7 | 51.8 | 227 | 749.1 | -70.1 | 1034 | 216 | 13.8 | 6.4 | 41 | 137 | -79.1 |
| Ca (%) | 39.59 | 43.04 | 0.72 | 1.68 | 2.17 | 7.17 | 8.69 | 21.68 | 24.1 | 1.87 | 7.77 | 5.62 | 18.54 | 11.17 | 26.87 | 30.51 | 0.18 | 0.6 | 0.55 | 1.82 | 13.57 |
| Mn (ppm) | 0.008 | 0.010 | 0.001 | 12.6 | 0.004 | 0.012 | 27.7 | 0.063 | 0.060 | 0.002 | 3 | 0.005 | 0.018 | -5.1 | 0.016 | 0.017 | 0.002 | 9.1 | 0.005 | 0.015 | 8.5 |
| Fe (ppm) | 0.032 | bd | - | - | - | - | - | 0.314 | 0.327 | 0.004 | 1.2 | 0.012 | 0.039 | 4.2 | 0.321 | 0.319 | 0.003 | 1 | 0.010 | 0.032 | -0.8 |
| Zn (ppm) | 6 | 8 | 3.2 | 37.7 | 9.5 | 31.4 | 40 | 66 | 74.2 | 2.2 | 2.9 | 6.5 | 21.6 | 12.4 | 49 | 53 | 1.7 | 3.3 | 5 | 17 | 7.1 |
| Sr (ppm) | 161 | 158 | 1.4 | 0.9 | 4.1 | 13.5 | -2.2 | 27 | 25.9 | 0.5 | 2.1 | 1.6 | 5.3 | -4.1 | 121 | 125 | 1 | 0.8 | 3 | 10 | 3.5 |
| Pb (ppm) | 3 | 8 | 1.5 | 19.7 | 4.6 | 15.2 | 160 | 27 | 32.7 | 1.7 | 5.1 | 5 | 16.6 | 21.1 | 4 | 7 | 1.6 | 22.6 | 5 | 16 | 81.3 |

|  | NCS DC 28202 |  |  |  |  |  |  | NCS DC 70306 |  |  |  |  |  |  |  | AR CaCO3 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | xCRM | xav | σ | RSD | LOD | LOQ | RD | xCRM | xav | σ | RSD | LOD | LOQ | RD | xRM |  | xav | σ | RSD | LOD | LOQ | RD |
| Mg (%) | 12.5 | 13.66 | 0.54 | 3.97 | 1.63 | 5.37 | 9.26 | 0.87 | 1.49 | 1.11 | 74.11 | 3.32 | 10.95 | 70.65 | 0.09 |  | bd | - | - | - | - | - |
| S (ppm) | 340 | 426 | 41.1 | 9.6 | 123 | 407 | 25.3 | 3924 | 2097 | 115.1 | 5.5 | 345 | 1140 | -46.6 | 699 |  | bd | - | - | - | - | - |
| Ca (%) | 22.01 | 25.07 | 0.26 | 1.06 | 0.79 | 2.62 | 13.9 | 34.42 | 38.41 | 1.89 | 4.91 | 5.66 | 18.67 | 11.59 | 39.1 |  | 44.18 | 0.9 | 2.04 | 2.7 | 8.9 | 13 |
| Mn (ppm) | 0.020 | 0.021 | 0.002 | 9.4 | 0.006 | 0.019 | 4.2 | 0.069 | 0.091 | 0.002 | 1.7 | 0.005 | 0.015 | 31.8 | 0.004 |  | bd | - | - | - | - | - |
| Fe (ppm) | 0.192 | 0.197 | 0.004 | 1.8 | 0.011 | 0.035 | 2.5 | 0.511 | 0.504 | 0.004 | 0.7 | 0.011 | 0.036 | -1.2 | 0.008 |  | bd | - | - | - | - | - |
| Zn (ppm) | 7 | 3 | 2.9 | 84.4 | 9 | 28 | -51.4 | 14 | 33 | 2.5 | 7.3 | 7 | 24 | 144.1 | 3 |  | 2 | 3.1 | 187.6 | 9 | 31 | -44.4 |
| Sr (ppm) | 72 | 73 | 1.1 | 1.6 | 3 | 11 | 1.5 | 477 | 477 | 4.2 | 0.9 | 13 | 42 | -0.1 | 414 |  | 431 | 2.6 | 0.6 | 8 | 26 | 4.2 |
| Pb (ppm) | 6 | 5 | 2.6 | 54.2 | 8 | 26 | -20 | 6 | 12 | 1.5 | 13.1 | 5 | 15 | 110.3 | 3 |  | 3 | 3.1 | 90.9 | 9 | 31 | 14.8 |

|  | BT5 |  |  |  |  |  |  | RM01 |  |  |  |  |  |  | RM04 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | xRM | xav | σ | RSD | LOD | LOQ | RD | xRM | xav | σ | RSD | LOD | LOQ | RD | xRM | xav | σ | RSD | LOD | LOQ | RD |
| Mg (%) | 0.54 | 2.23 | 0.88 | 39.67 | 2.65 | 8.85 | 310.9 | 0.17 | bd | - | - | - | - | - | 0.22 | 3.17 | 0.61 | 19.14 | 1.82 | 6.08 | 1326.3 |
| S (ppm) | 5903 | 2033 | 62.5 | 3.1 | 188 | 625 | -65.6 | 329 | bd | - | - | - | - | - | 1279 | bd | - | - | - | - | - |
| Ca (%) | 7.42 | 8.54 | 0.13 | 1.52 | 0.39 | 1.30 | 15.0 | 32.65 | 38.56 | 0.46 | 1.18 | 1.37 | 4.56 | 18.1 | 1.80 | 1.77 | 0.01 | 0.75 | 0.04 | 0.13 | -1.7 |
| Mn (ppm) | 22.573 | 23.340 | 0.085 | 0.37 | 0.256 | 0.852 | 3.4 | 0.762 | 0.798 | 0.012 | 1.56 | 0.037 | 0.124 | 4.7 | 5.306 | 4.928 | 0.062 | 1.26 | 0.187 | 0.622 | -7.1 |
| Fe (ppm) | 12.833 | 13.813 | 0.050 | 0.36 | 0.149 | 0.496 | 7.6 | 5.586 | 7.597 | 0.071 | 0.94 | 0.213 | 0.711 | 36.0 | 39.645 | 39.023 | 0.251 | 0.64 | 0.754 | 2.512 | -1.6 |
| Zn (ppm) | 14 | bd | - | - | - | - | - | 10 | 7 | 4.7 | 68.4236 | 14 | 47 | -31 | 46 | bd | - | - | - | - | - |
| Sr (ppm) | 161 | 142 | 2.7 | 1.9 | 8 | 27 | -11.7 | 148 | 151 | 2.8 | 1.8 | 8 | 28 | 2.1 | 123 | 91 | 2.3 | 2.5 | 7 | 23 | -26.1 |
| Pb (ppm) | 3 | bd | - | - | - | - | - | 3 | bd | - | - | - | - | - | 3 | 347 | 19.4 | 5.6 | 58 | 194 | 11456.7 |